qRT-PCR Protocol with Standard Curve

- Clean your bench and pipetman with RNase Away.

- Anywhere below that says H2O means DEPC H2O.

- For each gene you are testing, you will also do:

  - a standard curve using “bulk” RNA from either WT cells (29-13, e.g.) or -tet cells.

  - an H2O control containing primers, SYBR green, and H2O

  - a -RT control. Do an -RT for each sample (e.g., -tet and +tet), but it is OK to do it only on the housekeeping gene (e.g., β-tubulin) and not the gene of interest

DNase Treatment:  with Ambion DNaseI DNAfree kit

10 µg RNA												5 µl 10X DNaseI buffer											1 µl rDNaseI

0.5 µl RNAse inhibitor											H2O to 50 µl

Incubate at 37°C 30 min, add 1 µl rDNaseI, and incubate 37°C 1 hr.  Add 150 µl H2O followed by 200 µl phenol/chloroform/IAA (acidic pH). Spin 25 min at 4°C and take top layer. Precipitate RNA with 20 µl 3M NaOAc + 1 µl glycogen + 800 µl cold 100% EtOH in -20°C overnight.  Spin down RNA at 13,000rpm 30 min, remove supernatant, dry pellet 10 min, and resuspend pellet in 20 µl DEPC H2O.  Spec RNA.

Reverse Transcription Reaction: There are two different kits/protocols that can be used. Applied Biosystems Taqman kit (Brianna) or BioRad iScript kit (Kyle). Both are explained below.

With Applied Biosystems Taqman kit

4 µg RNA per reaction											5 µl Taqman RT buffer											11 µl 25mM MgCl2											10 µl dNTP mix												3.5 µl Random Hexamers										0.5 µl RNase Inhibitor											1.5 µl Multiscribe RT											H2O to 50 µl

Incubate at 25°C for 10 min, 48°C for 45 min, 95°C for 5 min, cool on ice, then add 1.5 µl RNaseH (Ambion/Applied Biosystems), incubate at 37°C for 20 min, then at 65°C for 20 min.

With BioRad iScript Select cDNA Synthesis kit

1 µg RNA per reaction

4 µl 5X iScript reaction mix

2 µl primer (oligo d(t)20, random hexamer, ect.)

1 µl iScript RT

DEPC H2O to 20 µl

Incubate at 42°C for 90 mins, 85°C for 5 mins, 12 for ∞

Q-PCR Set-up:

Assume a 1:1 conversion of RNA into cDNA. For example, if you started with 4 µg of RNA and ended up with 50 µl if cDNA, then you assume cDNA = 80ng/µl.

“Bulk” cDNA:

- RT dilute 80ng/uL at 1:5 = 16ng/ul

+ RT serial dilute 1X = 1:5 dilution from 80ng/uL = 16ng/uL, 								  1/4X = 1:4 dilution from 16ng/uL = 4ng/uL								  1/16X = 1:4 dilution from 4ng/uL = 1ng/uL								  1/64X = 1:4 dilution from 1ng/uL = 0.25ng/uL								  1/256X = 1:4 dilution from 0.25ng/uL = 0.0625ng/uL							  1/1024X = 1:4 dilution from 0.0625ng/uL = 0.015625ng/uL

Sample cDNA:

Dilution for gene of interest use 1:20 dilution of sample RT at 80ng/µl = 4ng/µl

Dilution for housekeeping (control) gene use 1:10 dilution of sample RT at 4ng/µl = 0.4ng/µl

Primers (all primer stocks at 100uM):

Add 15 µl of each forward and reverse primers to tube and dilute with 970 µl H2O, primer mix = 1.5 µM

Following plate map (example in Q-PCR plate map file) add 2.5 µL of cDNA to each well.  Mix 1.5 µM primers with SYBR green Master Mix in a tube (10 µl primers to 12.5 µl SYBR green per well for as many wells as needed) and add 22.5 µL of primer/SYBR mix to each well.  Cover plate with adhesive cover, use squeegee to press down cover, and place in machine.

Hit Read lab program, which should be:

Cycle 1, Step 1 = 50°C 2 min

Cycle 2, Step 1 = 95°C 10 min

Cycle 3 40X, Step 1 = 95°C 15 sec, Step 2 =60°C 1 min (take measurement)

Cycle 4 80X, Step 1 melt curve = 55°C 10 sec, then temp change in 0.5°C increments to 94.5°C (take meas.)